## Slide 1
2019 AWOP
National Meeting
Office of Ground Water and Drinking Water
Standards and Risk Management Division
Technical Support Center
August 7, 2019

## Slide 2: Drinking Water Methods: Evaluation and Approval
Recently Approved Methods – pH, online chlorine analysis and turbidity

## Slide 3: Drinking Water Methods
Methods are approved through a Regulation (Rule) or the Expedited Method Approval Process
Many methods developed at EPA or other government agencies
Some developed by vendors – e.g., Hach, Tintometer, Swan (turbidity)
Standards organizations - ASTM, Standard Methods
  EPA Drinking Water Alternate Test Procedure (ATP) Program – Evaluates new or modified test methods to determine if method is “equally effective” in performance relative to methods already approved in regulations

## Slide 4: EPA Drinking Water ATP Program
ATP program evaluates modified or new testing methods
  Accepted submissions are methods, not technologies
Methods must undergo sufficient validation to support their use at the national level (multi-lab validation/multi-DW matrices).
  Must demonstrate method robustness through method performance data and meeting QC acceptance criteria
  Single laboratory or Regional approvals apply to wastewater methods, not drinking water methods
ATP program does not have authority to approve alternate testing procedures

## Slide 5: Expedited Method Approval Process http://www.epa.gov/dwanalyticalmethods/expedited-drinking-water-analytical-method-approval-requirements
Alternate methods approved through a Federal Register Notice
  Once published in the FR, methods are fully U.S. EPA approved for regulatory compliance monitoring
Time required for approval significantly shortened
  Notice-and-comment rulemaking: 2-3 years following method evaluation
  Expedited process: < 1 year
  Expedited methods are in Appendix A to Subpart C of 40 CFR Part 141.
  State adoption of alternative test methods is optional; however, if these methods are used, laboratory certification requirements extend to the use of methods approved through the expedited process

## Slide 6: Expedited Method Approval Process http://www.epa.gov/dwanalyticalmethods/expedited-drinking-water-analytical-method-approval-requirements
Frequency of approvals
  Publishing FR notices on approximately an annual basis
Expedited Method Approvals could include:
  Methods evaluated through the EPA Drinking Water ATP program;
  Standard Methods and ASTM methods; and
  New or revised EPA methods

## Slide 7: Where can I find Approved Methods?
Approved methods are listed on OGWDW’s web site.
  http://www.epa.gov/dwanalyticalmethods/approved-drinking-water-analytical-methods
  Tables of approved methods
    Grouped by regulation/monitoring requirement, sorted by contaminant

## Slide 8: Method Focus: pH
Approved regulatory methods:
  EPA Method 150.1 (stand alone) and EPA Method 150.2 (continuous monitoring)
    Published in 1983
    Do not address newer sensor technologies available for pH determination
    Lack guidance on calibration frequency and verification

## Slide 9: Method Focus: pH
EPA Method 150.3:
  Single method addresses stand alone and continuous pH monitoring
  Describes calibration, calibration frequency and verification, and acceptance criteria
  Addresses direct (pH reference buffers) and indirect (grab sample comparison) calibration for continuous monitoring systems
  Provides guidance to assist operators in dealing with potential issues (e.g. effect of temperature on pH measurements
Approved for compliance monitoring in July 2017

## Slide 10: Method Focus: Online Free/Total Chlorine
EPA Method 334.0:
  Surface Water Treatment Rule (141.74(a)(2)) allowed the use of continuous monitoring without providing guidance
  EPA Method 334.0 establishes performance QC criteria for on-line analyzers
  On-line analyzers may be calibrated using aqueous standards or against grab sample measurements
  Weekly on-line analyzer calibration verification against grab sample measurement
  14-day initial demonstration of capability to confirm analyzer/grab sample agreement before using analyzer for compliance
Approved for compliance monitoring in November 2009

## Slide 11: Method Focus: Turbidity
Turbidity is a method-defined parameter:
  Not a specific chemical that is measured, but a physical parameter
  EPA Method 180.1 established design specifications, quality control and calibration requirements for turbidity measurement by nephelometry
  Primary standard (formazin) used for calibration
  Secondary standard used for routine daily calibration checks
Light sources
  EPA Method 180.1 set specifications for tungsten lamps
  Newer methods use LED or laser-based light sources

## Slide 12: Method Focus: Turbidity
Turbidity methods approved since 2009:
  Laser-based nephelometry
    Mitchell M5271, Hach Method 10258 (360o Nephelometry), Tintometer Lovibond 6000
  LED-based nephelometry
    Mitchell M5331, Swan AMI Turbiwell, Thermo AQ4500, Tintometer Lovibond 1000, Tintometer Lovibond 2000
Revised Hach 8195, Rev. 3.0 (tungsten/LED) & 10258 (Laser 360o)
  Previously – sealed standards could only be used as secondary calibration verification checks. Quarterly preparation of liquid standards were required
  Hach conducted long-term stability studies using sealed StabilCal for calib.
  The two revised methods were approved in October 2018, allowing the use of prefilled sealed StabilCal vials as primary calibration standards

## Slide 13: Questions?